# ACCEPTANCE TEST PROCEDURE

# FOR

# ILET-6-40-2 SYSTEM

# S/N: 25OR22244

# Contract: PO 1294

PROPRIETARY INFORMATION: The information contained in this document is proprietary to CI Systems and may not be reproduced, used, or disclosed without its prior written consent.

1.	General

1.1	This document describes the acceptance test procedures (ATP) for the ILET-6 major components and for the complete system.

1.2	The tests are defined and the detailed procedures are given in this document.

1.3	The required results and the actual system results are summarized in the acceptance test reports (ATR).

# 2.	Test Equipment

a.	Fiber optic light source CUDA MODEL 1.-150

b.	Autocollimator Nikon 6D/LUPI interferometer

c.	Diffusive transparent paper

d.	Fiber optic holder

e.	FLIR

f.          Dial gauge tester

# 3.	List of Tests

a.	Optical alignment test

b.	Filling the exit pupil test

c.	Mechanical integrity test

d.	Visual test

e.	Targets positioning accuracy

- Stray Light

- LOS Target check

- SR800N-4D Blackbody

- SR300N Integrating Sphere

NOTE:	The following tests must be performed before mounting the blackbody head onto the ILET-6 collimator.

# 4.	Test Procedure

4.1	Optical Alignment Test

4.1.1	Test by LUPI interferometer (internal procedure).

4.1.2	Set the autocollimator in front of the exit pupil of the collimator.

4.1.3	Set the Target wheel on a small four-bar target (if a four-bar target is not available set the Target wheel on a small round aperture).

4.1.4	Find the target in the autocollimator's field of view and, by using the autocollimator objective drum, obtain the sharpest image of the target.

4.1.5	Read the number in the autocollimator objective drum.  The result should be 100 ± 0.1%.

4.2       Filling the Exit Pupil Test

4.2.1	Position the autocollimator in front and centered to the aperture.

4.2.2	Put the diffusive transparent paper (e.g. wax paper) on the collimator exit pupil.

4.2.3	Use the fiber optic source to illuminate through an open aperture or through the largest pinhole target.

4.2.4	Check whether the distribution of light on the paper is centered and that it fills the complete exit pupil.

4.3	Mechanical Integrity Test

4.3.1	Turn the Target Wheel and make sure that it stops without backlash in every one of its positions.

4.3.2	Check the tightness of the screws and the general integrity of the instrument.

4.4	Visual Inspection

4.4.1	Verify target labels:

The 8 target positions should read the index numbers.

4.4.2  Check that the target number labels are clear and centered in the viewing 	window.

4.4.3	Check that the paint is uniform and unscratched.

4.4.4	Check the instrument label (model no., S/N, etc.) and that the label is well attached.

4.4.5	Check overall instrument cleanliness.

4.5	Target Positioning Repeatability & Accuracy

4.5.1	Position the autocollimator in front of the ILET-6 exit pupil.

4.5.2	Select the target with the smallest aperture.

4.5.3	Tilt the autocollimator until the target is accurately centered on the 				            autocollimator scale.

4.5.4    Target Repeatability - Turn the Target Wheel to a different position and back to the initial position. Measure the deviation in the target position, it should not be more than ± 0.2 mRad. Repeat this measurement 5 times.

4.5.5    Target Accuracy - Select the rest of the targets sequentially. On each target, measure the angular deviation from the center of the autocollimator scale.  The peak-to-peak position accuracy should be ± 0.2 mRad.

4.6	Stray light

4.6.1	Position the FLIR in front of the collimator

4.6.2    Rotate the target wheel and switch between all target positions.

4.6.3	Make sure no stray light is observed in the exit aperture of the collimator.

## 4.7       LOS alignment target

#### 4.7.1  	Perform a visual check of the LOS alignment target items: Target, LEDs, CCD camera, folding Mirror on rail.

#### 4.7.2 	 Attach LOS camera manufacturer COC

#### 4.7.3 	 Perform functionality test – using a reference mirror, make sure the light spot coming back from the target LED, is detectable and can be aligned to the center of the target cross.

#### 4.7.4 	Write Pass/Fail in the ATR document.

#### 4.8 	SR800N-4D Blackbody

#### 4.8.1 	Test Procedure:

#### 4.8.2	Verify that the ATP for the SR800N-4D is completed according to the procedure described in doc. # TA704-100-0010

## 4.9 	SR300N- Integrating Sphere

## 4.9.1 	Requirement:  ATP completion

4.9.2 	Test Procedure: # D120000005 Rev. C

## 4.9.3 	Verify that the ATP for the SR-300N is completed according to the procedure

## 4.9.4 	Write Pass/Fail in the ATR document Integrating Sphere

4.9.5    Verify all filters positions and check repeatability.

+ Add C.O.C documentation from supplier (for 2 filters)

ACCEPTANCE TEST REPORT   ILET-6-40-2

S/N: 25OR22244

The instrument tests performed according to the Acceptance Test Procedure for ILET-6 collimator.

The results are recorded in the following tables:

| TEST | REQUIRED RESULTS | ACTUAL RESULTS |
| --- | --- | --- |
| 1. Optical alignment | Autocollimator focusing ring reads 100 ± 0.1/ LUPI interferometer |  |
| 2. Filling exit pupil | Light beam is centered with exit pupil and fills it completely |  |
| 3. Mechanical check | The aperture wheel has 8 discrete positions. The wheel is positively held at each position. |  |
| 4. Visual check Paints, labels, cleanliness | Aperture wheel labels are according to target wheel drawings and requirements Paint is uniform, label well attached; instrument is free of dust and any kind of spurious particles. |  |
| 5. Target Positioning Accuracy | * Repeatability ±0.20mrad * Accuracy ±0.20mrad |  |
| 6. Stray Light | FLIR Inspection |  |

# 5.	Target Positioning Accuracy Test

| TARGET POSITION | ANGULAR ACCURACY [sec] Vertical Horizontal |
| --- | --- |
| 1. 4Bar 1.0 [Cy/mRad] |  |
| 2. 4Bar 1.5 [Cy/mRad] |  |
| 3. 4Bar 2.0 [Cy/mRad] |  |
| 4. 4Bar 2.5 [Cy/mRad] |  |
| 5. STEP HOR |  |
| 6. STEP VER |  |
| 7. SQUARE 1.5 [mRad] |  |
| 8. LOS Target |  |
| 9. 4Bar 0.15 [Cy/mRad] |  |
| 10. 4Bar 0.4 [Cy/mRad] |  |
| 11. 4Bar 0.6 [Cy/mRad] |  |
| 12. 4Bar 0.8 [Cy/mRad] |  |
| 13. USAF POSITIVE |  |
| 14. USAF NEGATIVE |  |
| 15. CROSS 0.5 [mRad] |  |

# The peak-to-peak angular accuracy is: 0.20mrad

Repeatability is better than 0.20mrad (with pinhole target).

Checked by:

Approved by:

NAME		    SIGNATURE 	    	        DATE