Acceptance Test Procedure (ATP)

For

METS-VS/S/L/VL-XX

S/N:  YYMETS-VS/S/L/VL-XXXXX-01

CONTRACT: XXXX

Proprietary Information:

The information contained in this document is proprietary to CI Systems and may not be reproduced, used, or disclosed without its prior written consent.

TABLE OF CONTENT

1.1	General	3

2	LIST OF TESTS	3

Table 1: List of Tests	3

2.1	List of Test Equipment	4

Table 2: List of Equipment	4

3	TEST PROCEDURES	4

3.1	Major Components	4

3.2	Visual Inspection and Verification:	5

3.3	Spectral Band	5

3.4	EFL & FOV:	5

3.5	Transmission:	6

3.6	Clear Aperture:	6

3.7	Target Wheel & Targets:	6

3.8	Boresight between Targets:	7

3.9	Cables	7

Table 1: Results Table	8

Table 2: Visual Inspection Results Table	9

Table 3: Laser Measurements (1 sets of 10 samples)	10

Table 4:  4-Bar Targets	10

Table 5:  Cross + Pinhole Targets	10

Table 6:  Special Targets	10

SCOPE

## General

- This document describes the Acceptance Test Procedure (ATP) for the METS collimator major components as well as for the integrated system.

- The tests are defined and the detailed procedures are given in this document.

- The system test results are summarized in the Acceptance Test Report (ATR).

- The METS Collimator performances are with respect to a standard laboratory environment:

| Environmental Temperature | : | 22 ±10ºC |
| --- | --- | --- |
| Relative Humidity (Non-Condensing) | : | 20% - 80%, non-condensing |
| Atmospheric pressure |  | 750+50mmHg/725-75mmHg |

# LIST OF TESTS

The ATP refers to the following tests:

Table 1: List of Tests

| Test No. | Test Name | Test Method | Notes |
| --- | --- | --- | --- |
|  | Major Components | Inspection |  |
|  | Visual inspection | Inspection |  |
|  | Spectral Band (Coating) |  | COC & COT |
|  | EFL & FOV |  | COC |
|  | Transmission (Coating) |  | COC & COT |
|  | Clear Aperture | Test |  |
|  | Target Wheel & Targets | Test | IAW Customer request |
|  | Boresight Between Targets | Test |  |

## List of Test Equipment

Table 2: List of Equipment

|  | Equipment | Manufacture | Model |
| --- | --- | --- | --- |
|  | LUPI Interferometer | CI-Systems |  |
|  | Laser | MeproLight | LRF16-52/54 |
|  | Profile Projector (for Target inspection) | Mitotoyo | PJ-300 |
|  | Roller & Caliber | Std. |  |
|  | Autocollimator | Nikon | 6D |
|  | FLIR | CEDIP | XXX |
|  | Dial Test Indicator | Mitutoyo |  |
|  | Photo meter | SpectiLight |  |

# TEST PROCEDURES

## Major Components

The METS collimator contains the following major components:

- METS Collimator Assembly.

- Motorized Target Wheel.

- IR Source: SR800N-XD-RR Blackbody + Controller.

- VIS Source: SR300N-X" integrating sphere + Controller.

- Motorized Source Stage.

- Motorized Focus stage.

- XY stage.

- LED Back Light.

- Target LOS with CCD&LED.

- LOS laser diode.

- Newport Photo Diode.

- Ophir Laser meter.

- Ophir power meter.

- Industrial PC.

- Gimbal.

- VRS.

### Test Procedures:

#### General: Inspect the METS system for its major component and verify existence of all quality assurance (QA) documents, as listed below.

#### Collimator:

#### Verify that the mirrors’ interferograms are attached.

#### Verify that the focusing interferogram is attached.

#### Visually inspect the collimator mechanical structure for defects.

#### Target Wheel & Targets:

#### Verify that the target wheel is rotated freely and is controlled via the controller.

#### Verify that the targets are set as described in the ATR document. (IAW customer request).

#### Cables:

#### Verify that all the cables are as described in the ATR document.

#### Write Pass/Fail in the ATR document.

## Visual Inspection and Verification:

### Test Procedures:

#### Inspect the following and verify its condition:

- External casing.

- Paint.

- Overall instrument cleanness.

- Cleanliness & Completeness of optical parts.

- Marking & Identification labels on system and each subsystem.

- Covering cap is suitable for the exit aperture.

- Exit aperture diameter of XX”.

- Stray Light-No stray light in the exit aperture of the collimator-Flir inspection.

- Target plates side towards the collimator emissivity of: 0.97  0.02 and side towards the blackbody high reflectivity.

- Targets:

  - Profile projector (PJ-300) inspection.

  - Mechanical integrity of all the targets (see target table) - supplied target inspection report.

  - Location of targets on the plates as defined in the target table in the ATR document.

#### Write Pass/Fail in the ATR document.

## Spectral Band

### Test Procedure:

#### Verify that the mirrors coating is according to required spectral band.

#### Attach manufacturer COC & COT.

#### Write Pass/Fail in the ATR document.

## EFL & FOV:

### Test Procedure:

#### Verify that the EFL is according to CI requirements.

#### Attach manufacturer COC.

#### Calculate the FOV using the target plate diameter and the EFL as in manufacturer            COT.

#### Record the test results and write Pass/Fail in the ATR document.

## Transmission:

### Test Procedure:

#### Verify that the mirror’s reflection is according to CI requirements.

#### Calculates the transmission using the reflection factor of both mirrors @  = 10.

#### Attached manufacturer COC.

#### Record the test results in the ATR document.

## Clear Aperture:

### Test Procedure:

#### Verify that the diameter of the parabolic mirror is according to CI requirements.

#### Attach manufacturer COC.

#### Use a ruler or any other measuring device and measure the collimator’s clear aperture diameter.

3.6.1.4 Record the test results in the ATR document.

## Target Wheel & Targets:

### Test Procedure:

#### Verify that the target wheel is running freely when motor power is disconnected.

#### Apply power and rotate the target wheel using the controller\computer, verify the wheel is running freely and stops according to command.

#### Verify that the targets are set on the wheel as required by system specifications.

#### Attach the targets dimension report.

#### Record the test results in the ATR document.

## Boresight between Targets:

### Test Procedure:

#### The boresight between targets should be  0.2 mrad.

#### Position an autocollimator in-front the METS aperture.

#### Focus the autocollimator until a sharp image is achieved.

#### Using the controller or computer set the target wheel to a chosen PH or Cross target.

#### Align the autocollimator until the center of the image is located at the center of the autocollimator’s cross reticule.

#### Rotate the target wheel to the next target.

#### Verify that the center of the target is located no more than  0.2 mRad

#### Repeat stage 3.8.1.6 & 3.8.1.7 for all the targets.

#### Record the test results in the ATR document.

## SR800N-XD black body

### Test Procedure:

#### Verify that the ATP for the SR-800N-XD is completed according to the procedure          described in doc. #TA7041000010.

#### Test Radiometric offset according to procedure D1506034250.

#### Record the test results in the ATR document.

## Motorized focus stage

### Test Procedure:

3.9.1.1 צריך להיות 3.17.1...Verify that finite distances of A[mm] & B[mm] are available in the SR800N controller.

3.9.1.2	Choose finite distance X[mm] and wait for the focus stage to rest in position.

Calculate the Dx shift of the stage: Dx[mm] = Fl2/(A-Fl)

3.9.1.3	Measure the distance of movement using a Poppy-test Dial Test Indicator. The result should be 9.32mm±0.5mm.

3.9.1.4	Choose finite distance Y[mm] and wait for the focus stage to rest in position.

Calculate the Dx shift of the stage: Dx[mm] = Fl2/(B-Fl)

3.9.1.5	Measure the distance of movement using a Poppy-test Dial Test Indicator. The result should be 47.2mm±0.9mm.

3.9.1.6	Record the test results in the ATR document.

## Cables

### Inspect the following and verify its condition

#### External casing.

#### Overall cleanness.

#### Marking & Identification labels on cable.

#### Write Pass/Fail in the ATR document.

Acceptance Test Results (ATR)

For

METS

S/N:  YYMETS-XX-

Table 1: Results Table

|  | Test name | Paragraph | Results | Pass/fail | Notes |
| --- | --- | --- | --- | --- | --- |
| 1 | Major Components | 3.1 | O.K. | PASS |  |
| 2 | Visual Inspection | 3.2 | O.K. | PASS | Refer To Table 2 |
| 3 | Wavelengths Region | 3.3 |  | PASS |  |
| 4 | EFL | 3.4 |  | PASS |  |
| 5 | FOV | 3.4 |  | PASS |  |
| 6 | Transmission | 3.5 |  | PASS | @10µ |
| 7 | Clear aperture | 3.6 |  | PASS |  |
| 8 | Target Wheel & Targets | 3.7 | O.K. | PASS |  |
| 9 | Boresight Between Targets | 3.8 | Accuracy | PASS | Repeatability |
| 10 |  | 3.9 | O.K. | PASS |  |
| 11 |  | 3.10 | O.K. | PASS |  |
| 12 |  | 3.11 | O.K. | PASS |  |

Table 2: Visual Inspection Results Table

| No. | Test name | Results | Pass/fail | Notes |
| --- | --- | --- | --- | --- |
| 1 | Visual Inspection | O.K. | PASS |  |
| 1.1 | External Casing | O.K. | PASS |  |
| 1.2 | Paint | O.K. | PASS |  |
| 1.3 | Overall Cleanliness | O.K. | PASS |  |
| 1.4 | Optics Cleanliness | O.K. | PASS |  |
| 1.5 | Markings & Identifications | O.K. | PASS |  |
| 1.6 | Covering Cup | O.K. | PASS |  |
| 1.7 | Exit Aperture Diameter |  | PASS |  |
| 1.8 | Reference Mirror | O.K. | PASS |  |
| 1.9 | Stray light | O.K. | PASS |  |
| 1.10 | Targets Plates | O.K. | PASS |  |
| 1.11 | Targets | O.K. | PASS |  |

- TARGETS

Note for Target Size calculation

Cycles/mRad and mRad units' calculations are based on a nominal effective focal length of XXXXmm.

Insert the Targets Table from the “Target Wheel” Drawing (xls file). Add a column with O.K.

Table 3: System Targets

Table 4: Laser Measurements (1 sets of 10 samples)

| With Lens [mJoul] | Direct Energy [mJoul] | Divided Energy [mJoul] | Notes |
| --- | --- | --- | --- |

Performed By:

| Name |  | Date |  | Signature |
| --- | --- | --- | --- | --- |

Approved By:

Software Eng.

| Name |  | Date |  | Signature |
| --- | --- | --- | --- | --- |

Project Eng.

| Name |  | Date |  | Signature |
| --- | --- | --- | --- | --- |

QA Manager:

| Name |  | Date |  | Signature |
| --- | --- | --- | --- | --- |